## MISD-USB1601
MISD(Modular Instrumentation Software Dictionary) is a list of methods, properties used in the C# programming. This list has both English and Chinese notations with JYTEK hardware. MISD was authored and published by JYTEK from 2019 to 2023. Copyright belongs to JYTEK. Please indicate the source when reprinting or using.(MISD模块仪器软件词典由上海简仪科技有限公司撰写并发布，2019-2023，版权归简仪科技，转载或使用时请标明出处。)
Class | Function Group | Public Symbol | Public Classs, Method, Property etc | Type | 中文注释 | English Notation | 1601
AI | AITask | JYPXIeHardwareAITask | JYPXIeHardwareAITask(int slotNum) | Constructor | 构造函数。用来构造一个AI任务对象，slotNum板卡槽位号。硬件的唯一标志。 | Constructor. Used to create an AI task object. | Y
AI | AITask | AddChannel | AddChannel(int[] chnsId,double[] rangeLow,double[] rangeHigh,terminal,coupling,enableIEPE) | Method | 添加一个通道组。参数chnsID:一组通道物理序号; rangeLow:一组通道量程下限; rangeHigh:一组通道量程上限; terminal:端口输入模式配置,RSE/NRSE/Differential/Pseudo-differential；coupling耦合：由enum AC/DC 硬件选; enableIEPE：True/False启用/禁用IEPE激励。 | Add a group of channels. Parameters chnsID:a group of channel physical numbers; rangeLow:Lower range limit for a group of channels; rangeHigh:Higher range limit for a group of channels; terminal: terminal input mode configuration；coupling: enum AC/DC hardware specific; enableIEPE:True/false to enable or disable IEPE excitation, hardware specific. | Y
AI | AITask | AddChannel Reload | AddChannel(int[] chnsId,double rangeLow,double rangeHigh,terminal,coupling,enableIEPE) | Method | 添加一个通道组。参数chnsID:一组通道物理序号; rangeLow:公用通道量程下限; rangeHigh:公用通道量程上限; terminal:端口输入模式配置,RSE/NRSE/Differential/Pseudo-differential；coupling耦合：由enum AC/DC 硬件选; enableIEPE：True/False启用/禁用IEPE激励。 | Add a group of channels. Parameters chnsID:a group of channel physical numbers; rangeLow:common lower range limit for the group; rangeHigh: common higher range limit for the group; terminal: terminal input mode configuration；coupling: enum AC/DC hardware specific; enableIEPE:True/false to enable or disable IEPE excitation, hardware specific. | Y
AI | AITask | AddChannel Reload | AddChannel(int chnId,double rangeLow,double rangeHigh,terminal,coupling,enableIEPE) | Method | 添加一个通道。参数chnID:一个通道物理序号; rangeLow:通道量程下限; rangeHigh:通道量程上限; terminal:端口输入模式配置,RSE/NRSE/Differential/Pseudo-differential；coupling耦合：由enum AC/DC 硬件选; enableIEPE：True/False启用/禁用IEPE激励。 | Add one channel. Parameters chnID: channel's physical number; rangeLow: lower range limit for the channel; rangeHigh: higher range limit for the channel; terminal: terminal input mode configuration；coupling: enum AC/DC hardware specific; enableIEPE:True/false to enable or disable IEPE excitation, hardware specific. | Y
AI | AITask | GetRecordPreviewData | GetRecordPreviewData(ref double[,] buf, int samplesPerChannel, int timeout) | Method | 获取流盘时预览的二维数据，每一列代表一个通道。参数 buf:用户定义返回数据的缓冲区; samplesPerChannel:每通道要预览的数据;timeout:超时时间，单位ms，-1为持续等待直至完成。 | Preview two dimentional data when streaming is in progress. Each column represents one channel. Parameters buf: user specified buffer for preview data; samplesPerChannel: number of preview samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | GetRecordPreviewData Reload | GetRecordPreviewData(ref double[,] buf, int timeout = -1) | Method | 获取流盘时预览的二维数据，每一列代表一个通道，每一通道的数据个数是由缓冲区大小来计算的。参数 buf:用户定义返回数据的缓冲区; samplesPerChannel:每通道要预览的数据;timeout:超时时间，单位ms，-1为持续等待直至完成。 | Preview two dimentional data when streaming is in progress. Each column represents one channel. Samples per channel is caculated from the buffer size. Parameters buf: user specified buffer for preview data; samplesPerChannel: number of preview samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | GetRecordPreviewData Reload | GetRecordPreviewData(ref double[] buf, int samplesPerChannel, int timeout) | Method | 获取流盘时预览的单通道数据。参数 buf:用户定义返回数据的缓冲区; samplesPerChannel:每通道要预览的数据;timeout:超时时间，单位ms，-1为持续等待直至完成。 | Preview one channel data when streaming is in progress. Parameters buf: user specified buffer for preview data; samplesPerChannel: number of preview samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | GetRecordPreviewData Reload | GetRecordPreviewData(ref double[] buf, int timeout = -1) | Method | 获取流盘时预览的单通道数据，通道的数据个数是由缓冲区大小来计算的。参数 buf:用户定义返回数据的缓冲区; timeout:超时时间，单位ms，-1为持续等待直至完成。 | Preview one channel data when streaming is in progress. Samples per channel is caculated from the buffer size. Parameters buf: user specified buffer for preview data; timeout: timeout in mS before generating error message, timeout =-2 keep waiting unti | Y
AI | AITask | GetRecordStatus | GetRecordStatus(out double recordedLength, out bool recordDone) | Method | 获取预览状态。参数recordedLength:已流盘的长度; recordDone:流盘是否结束。 | Get the streaming status.Parameters recordLength:Length of the streamed. recordDone = false :streaming is not over.recordDone = true :streaming is over. | Y
AI | AITask | ReadData | ReadData(ref double[,] buf, int samplesPerChannel, int timeout) | Method | 按列读取采集到的多通道数据，存入缓冲区，一列代表一个通道的数据。参数 buf:缓冲区；samplesPerChannel：每通道读取的数据；timeout mS 最长等待时间，-1持续等待直至完成。 | Read the multi-channel data by columns and store it in a data buffer. Each channel is represented by one column of data. Parameters buf:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadData Reload | ReadData(ref double[,] buf, int timeout) | Method | 按列读取采集到的多通道数据，存入缓冲区，一列代表一个通道的数据，每通道的数据个数samplesPerChannel是根据缓冲区大小算出来的。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成。 | Read the multi-channel data by columns and store it in a data buffer. Each channel is represented by one column of data. samplePerChannel is computed by the buffer size. Parameters buf:data buffer;timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadData Reload | ReadData(ref double[] buf, int samplesPerChannel, int timeout) | Method | 读取采集到的单通道数据，存入缓冲区。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成。 | Read the one channel data and store it in a data buffer. Parameters buf:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadData Reload | ReadData(ref double[] buf, int timeout) | Method | 读取采集到的单通道数据，存入缓冲区。单通道点数由缓冲区长度计算。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成。 | Read One channel data and store it in a data buffer. samplePerChannel is computed by the buffer size. Parameters buf:data buffer;timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadData Reload | ReadData(IntPtr buf, int samplesPerChannel, int timeout) | Method | 读取采集到的单一或多通道电压值，存入缓冲区。缓冲区由C#数据类型IntPtr所指。参数 buf：指向缓冲区的指针；samplesPerChannel：每通道读取的数据；timeout mS 最长等待时间，-1持续等待直至完成。。只适用于一些特殊场合如与其它软件配合时使用。 | Read one or multi-channel data by columns and store it in a data buffer. The buffer is pointed by C# data type IntPtr. Parameters buf:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. Only used for very special occations such as interface with other software. | Y
AI | AITask | ReadRawData | ReadRawData(ref int[,] buf, int samplesPerChannel, int timeout) | Method | 按列读取采集到多通道的原始数据，存入缓冲区，一列代表一个通道的数据，每通道的数据个数samplesPerChannel是根据缓冲区大小算出来的。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成。 | Read the multi-channel raw data by columns and store it in a data buffer. Each channel is represented by one column of data. Parameters buf:data buffer;timeout: samplesPerChannel:number of samples per channel; timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadRawData Reload | ReadRawData(ref int[,] buf, int timeout) | Method | 按列读取采集到的多通道原始数据，存入缓冲区，一列代表一个通道的数据，每通道的数据个数samplesPerChannel是根据缓冲区大小算出来的。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成。 | Read the multi-channel raw data by columns and store it in a data buffer. Each channel is represented by one column of data. samplePerChannel is computed by the buffer size. Parameters buf:data buffer;timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadRawData Reload | ReadRawData(ref int[] buf, int samplesPerChannel, int timeout) | Method | 读取采集到的单通道原始数据，存入缓冲区。参数 buf:缓冲区；samplesPerChannel：每通道读取的数据;timeout mS 最长等待时间，-1持续等待直至完成。 | Read the one channel raw data and store it in a data buffer. Parameters buf:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadRawData Reload | ReadRawData(ref int[] buf, int timeout) | Method | 读取采集到的单通道原始数据，存入缓冲区。单通道点数由缓冲区长度计算。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成。 | Read One channel raw data and store it in a data buffer. samplePerChannel is computed by the buffer size. Parameters buf:data buffer;timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. | Y
AI | AITask | ReadRawData Reload | ReadRawData(IntPtr buf, int samplesPerChannel, int timeout) | Method | 读取单一或多通道的原始数据，缓冲区由C#数据类型IntPtr所指。参数 buf：指向缓冲区的指针。只适用于一些特殊场合如与其它软件配合时使用。 | Read one or multi-channel raw data by columns and store it in a data buffer. The buffer is pointed by C# data type IntPtr. Parameters buf:data buffer;timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. Only used for very special occations such as interface with other software. | Y
AI | AITask | DetectOpenThermocouple | DetectOpenThermocouple(BurnoutCurrentValue currentValue) | Method | 开路检测 | Perform open thermocouple detection (OTD) to determine if there are any enabled channels do not have a thermocouple connectecd to the modul | NA
AI | AITask | SetCJTemperature | SetCJTemperature(int chnId, double temperature) | Method | 设置指定通道冷端补偿温度 | Set cold junction (reference junction) temperature with a specified channel | NA
AI | AITask | SetCJTemperature Reload | SetCJTemperature(double[] temperature) | Method | 设置所有通道冷端补偿温度 | Set cold junction (reference junction) temperature to all channels that has been added | NA
AI | AITask | ReadSinglePoint | ReadSinglePoint(ref double[] buffer) | Method | 非缓冲式读取每个通道的一个样点。参数 buf: 读取的电压值。缓冲区必须大于通道数。 | Read one sample for all channels in unbuffered reading. Parameter buf: values read. Buffer size >= number of channels. | Y
AI | AITask | ReadSinglePoint Reload | ReadSinglePoint(ref double readValue, int channel) | Method | 非缓冲式读取指定通道的一个样点。参数 readValue: 读取的电压值；channel:指定的通道。 | Read one sample for specified channel in unbuffered reading. Parameters readValue: value read; channel:specified channel numbers. | Y
AI | AITask | RemoveChannel | RemoveChannel(int chnId) | Method | 删除指定通道号的通道,为-1则删除所有通道。参数：chnID:通道号。 | Delete the channel with the specified channel number. If it is -1, delete all channels. Parameter chnID:specified channel ID. | Y
AI | AITask | SendSoftwareTrigger | SendSoftwareTrigger() | Method | 产生一个软件触发信号。 | Generate a software trigger. | Y
AI | AITask | WaitUntilDone | WaitUntilDone(timeout=-1) | Method | 等待当前任务完成。参数 timeOut: timeOut=整数N，等待N毫秒，如果没有完成，报错，timeOut=-1, 等待直至完成。 | Waiting until current task to complete. Parameter timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | Y
AI | AITask | Start | Start() | Method | 启动AI任务 | Start AI task | Y
AI | AITask | Commit | Commit() | Method | 提交AI任务 | Commit AI task | NA
AI | AITask | Stop | Stop() | Method | 停止AI任务 | Stop AI task | Y
AI | AITask | AvailableSamples | AvailableSamples | Property | 缓冲区内可以读取的每通道点数 | Number of points per channel that can be read in the buffer | Y
AI | AITask | Advanced | Advanced | Property | 高级属性 | Advanced properties | NA
AI | AITask | BufLenInSamples | BufLenInSamples | Property | 缓冲区最大能容纳的每通道点数 | The maximum number of points per channel that the buffer can hold | Y
AI | AITask | TransferedSamples | TransferedSamples | Property | 已经传输的数据量 | The number of samples that have been transfered from the local buffer | Y
AI | AITask | Channels | Channels | Property | 通道列表,通道数随硬件变化。 | List of channels. Number of channels varies with hardware. | Y
AI | AITask | ClockSource | ClockSource | Property | 时钟源属性。enum External/Internal。 | Clock edge. enum Rising/Falling. | NA
AI | AITask | ClockEdge | ClockEdge | Property | 时钟边沿属性。enum Rising/Falling。 | Clock Source. enum External/Internal. | NA
AI | AITask | DisbaleCalibration | DisbaleCalibration | Property | 禁用校准系数 | Disalble calibration. | Y
AI | AITask | Mode | Mode | Property | AI采集模式：单点采样，连续，有限点采样，流盘。enum Single/Finite/Continuous/Record | AI acquisition mode：enum Single/Finite/Continuous/Record. | Y
AI | AITask | SampleClock | SampleClock | Property | 采样时钟配置。 | Sample Clock | Y
AI | AITask | SignalExport | SignalExport | Property | 信号路由 | Signal routing | Y
AI | AITask | AISync | Sync | Property | 同步参数配置 | Sync configuration | NA
AI | AITask | SampleRate | SampleRate | Property | 每通道采样率 | Sampling rate per channel. | Y
AI | AITask | SamplesToAcquire | SamplesToAcquire | Property | 有限点采集模式下每通道需要读取的采样点数 | Number of sampling points to be read per channel in Finite mode only. | Y
AI | AITask | Record | Record | Property | 流盘相关参数设置 | Data streaming related parameter settings | Y
AI | AITask | Trigger | Trigger | Property | 触发属性。由AITrigger()配置。 | AI trigger property configured by AI Trigger(). | Y
AI | AITask | JYXXDevice | Device | Property | 设备属性配置 | Device property set | Y
AI | AITask | BuildInCJCEnabled | BuildInCJCEnabled | Property | 冷端补偿使能 | Build-in cold junction compensation enabled. | NA
AI | AITask | AITimingMode | TimingMode | Property | 时钟等级模式 | Timing mode, can be set to automatic, or manually selected as level 0 to level 5. | NA
AI | AITask | AIChannelTopology | ChannelTopology | Property | 通道拓扑结构 | The channel topology currently in use. 2-wire and 3-wire RTD connection is supported in 32-channel mode, and 2-wire, 3-wire, and 4-wire RTD connection is supported in 20-channel mode. | NA
AI | AIAdvanced | AIAdvanced | AIAdvanced() | Constructor | 构造函数。用来构造AI高级设置 | Constructor to create an advanced AI task object. | NA
AI | AIAdvanced | AIAdvanced | ConvertRate | Property | AI高级设置:转换速率。默认由系统设置为SampleRate*ChannelCount。用户可以设置更大的扫描速率。这样扫描采集会在一次采样间隔的前期完成，可以降低通道之间采集的间隔时间。 | AI Advanced Setting: ConvertRate，The default is set by the system to SampleRate*ChannelCount. The user can set a higher scan rate. In this way, the scanning collection will be completed in the early period of a sampling interval, which can reduce the collection interval between channels. | NA
AI | AIAdvanced | AIAdvanced | DSMode | Property | 设置DSMode | config DSMode | NA
AI | AITrigger | AITrigger | AITrigger() | Constructor | 构造函数，创建AITrigger任务对象。用来构造AI Trigger | Constructor to create an AITrigger task object. Used to build AI Trigger | Y
AI | AITrigger | AITrigger | Analog | Property | 模拟触发属性。由三个部分组成。1，Condition enum AboveLevel/BelowLeve/RisingHysteresis/FallingHysteresis/EnteringWindow/LeavingWindow。和其它硬件相关; 2，Level enum Level/HighLevel/LowLevel; 3，Source: 硬件相关enum。 | AITrigger Analog has three parts. 1 Condition: enum AboveLevel/BelowLevel/RisingHysteresis/FallingHysteresis/EnteringWindow/LeavingWindow。; 2: Level: enum Level/HighLevel/LowLevel; 3: Source:enum hardware specific. | NA
AI | AITrigger | AITrigger | Delay | Property | AI触发被延迟了AI timebase Ticks数。 | Number of timebase Ticks delayed before AI triggering. | NA
AI | AITrigger | AITrigger | Digital | Property | AI数字触发.由三个属性组成。1，Source 由硬件定；2，Edge enum Rising/Falling；3，Gate enum HighActive/LowActive, 高电平/低电平触发。 | AITrigger Digital has three properties. 1: Source, hardware specific; 2: Edge enum Rising/Falling; 3: Gate enum HighActive/LowActive, high and low gate level triggering. | Y
AI | AITrigger | AITrigger | ReTriggerCount | Property | AI重触发次数。为0时不重复触发，>0时为重复触发次数,为-1时无限重触发。 | Number of repeat triggers. No repeat trigger when 0, or the number of repeated triggers. Infinite trigger when -1. | NA
AI | AITrigger | AITrigger | Type | Property | AI触发类型。enum Analog/Digital/Immediate/Soft。 | AI trigger type. enum Analog/Digital/Immediate/Soft. | Y
AI | AITrigger | AITrigger | Mode | Property | AI触发模式。由enum Start/Reference/Pause定。 | AI trigger mode. enum Start/Reference/Pause. | NA
AI | AITrigger | AITrigger | PreTriggerSamples | Property | AI预触发点数。 | AI pretrigger points. | NA
AI | AIRecord | AIRecord | AIRecord() | Constructor | 构造函数。用来构造流盘任务任务对象。 | Constructor to create a streaming task object. | Y
AI | AIRecord | AIRecord | FilePath | Property | 用户输入文件路径 | User specified file path. | Y
AI | AIRecord | AIRecord | FileFormat | Property | 文件格式.enum BIN。 | Streaming file format. enum BIN. | Y
AI | AIRecord | AIRecord | Length | Property | 用户指定流盘长度（秒），仅用于有限，Finite，模式。 | User specified stream length in second. Only valid in the Finite mode. | Y
AI | AIRecord | AIRecord | Mode | Property | 流盘模式。enum Finite/Infinite。 | Two stream mode from enum Finite/Infinite. | Y
AI | AIChannel | AIChannel | AIChannel() | Constructor | 构造函数，用来创建AIChnParam任务对象。 | Constructor to create an AIChnParam task object. | Y
AI | AIChannel | AIChannel | ChannelID | Property | 物理通道序号 | AI channel ID. | Y
AI | AIChannel | AIChannel | RangeHigh | Property | 通道量程上限 | High range limit an AI channel. | Y
AI | AIChannel | AIChannel | RangeLow | Property | 通道量程下限 | Low range limit an AI channel. | Y
AI | AIChannel | AIChannel | Terminal | Property | 端口输入模式配置,RSE/NRSE/Differential/Pseudodifferential | Configure AI terminal. Normally RSE/NRSE/Differential/Pseudodifferential. Hardware specific. | NA
AI | AIChannel | AIChannel | Coupling | Property | 通道耦合方式，AC/DC | Coupling Mode. Normally AC/DC. | NA
AI | AIChannel | AIChannel | ThemocoupleType | Property | 热电偶类型 | Thermocouple type | NA
AI | AIChannel | AIChannel | CustomizedCJTemperature | Property | 冷端补偿温度 | Customized cold junction (reference junction) temperature (℃) for compensation | NA
AI | AIChannel | AIChannel | MaxVoltageRange | Property | 最大电压范围 | Voltage range upper limit in V | NA
AI | AIChannel | AIChannel | MaxTemperatureRange | Property | 最大温度范围 | Temperature range upper limit(Ω) of the channel. | NA
AI | AIChannel | AIChannel | MaxResistanceRange | Property | 最大电阻范围 | Resistance range upper limit(Ω) of the channel. | NA
AI | AIAnalogTrigger | AIAnalogTrigger | AIAnalogTrigger() | Constructor | 构造函数，创建AITrigger任务对象。 | Constructor to create an AITrigger task object. Used to build AI Trigger | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Comparator | Property | 模拟比较器，enum Edge/Hysteresis/Window | Analog Comparator,enum Edge/Hysteresis/Window | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Edge | Property | 模拟边沿比较器 | Analog edge comparator | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Hysteresis | Property | 模拟迟滞比较器 | Analog hysteresis comparator | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Window | Property | 模拟窗比较器 | Analog window comparator | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Condition | Property | 设置模拟触发条件。触发值由HighLevel，LowLevel共同决定。enum RisingHysteresis/FallingHysteresis/EnteringWindow/LeavingWindow。RisingHysteresis:触发信号高于HighLevel;FallingHysteresis:触发信号低于LowLevel; EnterWindow: 触发信号进入LowLevel and HighLevel区域; LeavingWindow：触发信号离开LowLevel and HighLevel区域. | Configure analog trigger condition. TriggerValue is determined by properties HighLevel and LowLevel. Enum RisingHysteresis/FallingHysteresis/EnteringWindow/LeavingWindow. RisingHysteresis:triggering when above HighLevel;FallingHysteresis:triggering when below LowLevel; EnterWindow: triggering when the value fall between LowLevel and HighLevel; LeavingWindow：triggering when teh value falls outside LowLevel and HighLevel. | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Level | Property | 设置电平触发条件。 | Set level triggering value used in Condition triggering. | NA
AI | AIAnalogTrigger | AIAnalogTrigger | Source | Property | 设置模拟触发源。 | Set analog trigger source | NA
AI | AIAnalogTrigger | AIAnalogTrigger | HighLevel | Property | 设置高电平触发条件。 | Set high level triggering value used in Condition triggering. | NA
AI | AIAnalogTrigger | AIAnalogTrigger | LowLevel | Property | 设置低电平触发条件。 | Set low low level triggering value used in Condition triggering. | NA
AI | AIAnalogTrigger | AIAnalogEdgeComparator | AIAnalogEdgeComparator() | Constructor | 构造函数，创建AIAnalogEdgeComparator任务对象。 | Constructor to create an AITrigger task object. Used to build AIAnalogEdgeComparator | NA
AI | AIAnalogTrigger | AIAnalogEdgeComparator | Slope | Property | 设置边沿触发条件enum rising/falling | Config Edge trigger condition， enum rising/falling | NA
AI | AIAnalogTrigger | AIAnalogEdgeComparator | Threshold | Property | 设置边沿触发阈值 | Config Edge trigger threshold | NA
AI | AIAnalogTrigger | AIAnalogHysteresisComparator | AIAnalogHysteresisComparator() | Constructor | 构造函数，创建AIAnalogHysteresisComparator任务对象。 | Constructor to create an AITrigger task object. Used to build AIAnalogHysteresisComparator | NA
AI | AIAnalogTrigger | AIAnalogHysteresisComparator | Slope | Property | 设置迟滞触发条件enum rising/falling | Config Hysteresis trigger condition， enum rising/falling | NA
AI | AIAnalogTrigger | AIAnalogHysteresisComparator | HighThreshold | Property | 设置迟滞触发高阈值 | Config Hysteresis trigger high threshold | NA
| AIAnalogTrigger | AIAnalogHysteresisComparator | LowThreshold | Property | 设置迟滞触发低阈值 | Config Hysteresis trigger low threshold | NA
AI | AIAnalogTrigger | AIAnalogWindowComparator | AIAnalogWindowComparator() | Constructor | 构造函数，创建AIAnalogWindowComparator任务对象。 | Constructor to create an AITrigger task object. Used to build AIAnalogWindowComparator | NA
AI | AIAnalogTrigger | AIAnalogWindowComparator | Condition | Property | 设置模拟窗触发条件enum EnteringWindow/LeavingWindow | Config Window trigger condition， enum EnteringWindow/LeavingWindow | NA
AI | AIAnalogTrigger | AIAnalogWindowComparator | HighThreshold | Property | 设置模拟窗触发高阈值 | Config Window trigger high threshold | NA
| AIAnalogTrigger | AIAnalogWindowComparator | LowThreshold | Property | 设置模拟窗触发低阈值 | Config Window trigger low threshold | NA
AI | AIDigitalTrigger | AIDigitalTrigger | AIDigitalTrigger() | Constructor | 构造函数，创建AIDigitalTrigger任务对象。 | Constructor to create an AITrigger task object. Used to build AI Trigger | Y
AI | AIDigitalTrigger | AIDigitalTrigger | Edge | Property | 获取或设置数字触发条件。enum Rising/Falling。 | Get or set analog trigger condition.Possible enum AboveLevel/BelowLevel. AboveLevel:triggering above the set level. BelowLevel:triggering below the set level. | Y
AI | AIDigitalTrigger | AIDigitalTrigger | Source | Property | 获取或设置数字触发源。硬件相关。 | Set triggering source. Hardware spcecific | Y
AI | AIDigitalTrigger | AIDigitalTrigger | Gate | Property | 设置数字电平触发类型：enum HighActive/LowActive. | Configure digital level trigger type : enum HighActive/LowActive. HighActive: triggering when source is high active; LowActive: triggering when source is low active. | NA
AI | AISampleClock | AISampleClock | AISampleClock() | Constructor | 构造函数，创建AISampleClock任务对象。 | Constructor to create an AISampleClock object. | Y
AI | AISampleClock | AISampleClock | Source | Property | 采样时钟源，enum Internal/External | SampleClockSource，enum Internal/External | Y
AI | AISampleClock | AISampleClock | External | Property | 外部采样时钟 | External Sample Clock | Y
AI | AIExternalSampleClock | AIExternalSampleClock | AIExternalSampleClock() | Constructor | 构造函数，创建AIExternalSampleClock任务对象。 | Constructor to create an AIExternalSampleClock object. | Y
AI | AIExternalSampleClock | AIExternalSampleClock | Terminal | Property | 外部时钟输入端口 | External sample clock input terminal | Y
AI | AIExternalSampleClock | AIExternalSampleClock | ExpectedRate | Property | 外部时钟预期的频率 | External sample clock expected rate | Y
AI | AISignalExport | AISignalExport | AISignalExport() | Constructor | 构造函数，创建AISignalExport任务对象。 | Constructor to create an AISignalExport object. | Y
AI | AISignalExport | AISignalExport | Add(AISignalExportSource source, SignalExportDestination destination) | Method | 添加一路信号路由，将源信号路由至一个目的端口 | Add a one signal route, route the source signal to a destination terminal | Y
AI | AISignalExport | AISignalExport | Add(AISignalExportSource source, List<SignalExportDestination> destinations) | Method | 添加一路信号路由，将源信号路由至多个目的端口 | Add a one signal route, route the source signal to multiple destination terminals | Y
AI | AISignalExport | AISignalExport | Clear( SignalExportDestination destination) | Method | 删除一路信号路由，将源信号与目的端口断开 | Removes a signal route, disconnecting the source signal from the destination terminal | Y
AI | AISignalExport | AISignalExport | ClearAll() | Method | 删除所有存在的信号路由 | Delete all existing signal routes | Y
AI | AISync | AISync | AISync() | Constructor | 构造函数，创建AISync任务对象。 | Constructor to create an AISync object. | NA
AI | AISync | AISync | Topology | Property | 同步拓扑结构设置 | Sync Topology Configuration | NA
AI | AISync | AISync | TriggerRouting | Property | 同步触发信号路由 | Sync trigger signal routing terminal | NA
AI | AISync | AISync | PulseRouting | Property | 同步脉冲信号路由 | Sync pulse signal routing terminal | NA
AO | AOTask | JYHardwareAOTask | JYHardwareAOTask(int slotNum) | Constructor | 构造函数。用来建立AO任务对象。slotNum板卡槽位号。硬件的唯一标志。 | Constructor. Used to create AO task object. | Y
AO | AOTask | AddChannel | AddChannel(int[] chnsId, double[] rangeLow, double[] rangeHigh) | Method | 添加多个AO通道 | Add multiple AO channels | NA
AO | AOTask | AddChannel Reload | AddChannel(int[] chnsId, double rangeLow, double rangeHigh) | Method | 添加多个AO通道 | one or a group of channel physical numbers | Y
AO | AOTask | AddChannel Reload | AddChannel(int chnId, double rangeLow = -10, double rangeHigh = 10) | Method | 添加单通道(chnId=-1时添加所有通道) | Add a single channel (add all channels when chnId=-1) | Y
AO | AOTask | RemoveChannel | RemoveChannel(int chnId) | Method | 删除指定通道号的通道，(chnId=-1时删除所有通道) | Delete the channel with the specified channel number. (When chnId=-1, delete all channels) | Y
AO | AOTask | SendSoftwareTrigger | SendSoftwareTrigger() | Method | 产生一个软件触发信号。 | Generate a software trigger. | Y
AO | AOTask | Start | Start() | Method | 开始 | Start AO Task | Y
AO | AOTask | Stop | Stop() | Method | 停止 | Stop AO Task | Y
AO | AOTask | WaitUntilDone | WaitUntilDone(timeout=-1) | Method | 等待当前任务完成。参数 timeOut: timeOut=整数N，等待N毫秒，如果没有完成，报错，timeOut=-1, 等待直至完成。 | Waiting until current task to complete. Parameter timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | Y
AO | AOTask | WriteData | WriteData(double[,] buf, int timeout = 10000) | Method | 按列向AO缓冲区写入多通道数据，一列代表一个通道的数据。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成，预置值10000mS。 | WriteData multi-channel data to AO buffer, Each column represents one channel. Parameters. buf:AO buffer for AO data;timeout:timeout in mS, default 10000mS. | Y
AO | AOTask | WriteData Reload | WriteData(double[] buf, int timeout = 10000) | Method | 按列向AO缓冲区写入单一通道数据。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成，预置值10000mS。 | WriteData to one channel AO buffer.Parameters buf:AO buffer for AO data;timeout:timeout in mS, default 10000mS. | Y
AO | AOTask | WriteData Reload | WriteData(IntPtr buf, int samplesPerChannel, int timeout = 10000) | Method | 向AO缓冲区写入单一或多通道数值。缓冲区由C#数据类型IntPtr所指。参数 buf：指向缓冲区的指针；samplesPerChannel：每通道写入的数据；timeout mS 最长等待时间，-1持续等待直至完成。。只适用于一些特殊场合如与其它软件配合时使用。 | Write to AO buffer one or multi-channel data. The buffer is pointed by C# data type IntPtr. Parameters buf:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. Only used for very special occations such as interface with other software. | Y
AO | AOTask | WriteRawData | WriteRawData(ushort[,] buf, int timeout = 10000) | Method | 按列向AO缓冲区写入多通道原始数据，一列代表一个通道的数据。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成，预置值10000mS。 | WriteData multi-channel raw data to AO buffer, Each column represents one channel. Parameters. buf:AO buffer for AO data;timeout:timeout in mS, default 10000mS. | NA
AO | AOTask | WriteRawData Reload | WriteRawData(ushort[] buf, int timeout = 10000) | Method | 按列向AO缓冲区写入单一通道数据。参数 buf:缓冲区；timeout mS 最长等待时间，-1持续等待直至完成，预置值10000mS。 | WriteData to one channel AO buffer.Parameters buf:AO buffer for AO data;timeout:timeout in mS, default 10000mS.
AO | AOTask | WriteRawData Reload | WriteRawData(IntPtr buf, int samplesPerChannel, int timeout = 10000) | Method | 向AO缓冲区写入单一或多通道原始数值。缓冲区由C#数据类型IntPtr所指。参数 buf：指向缓冲区的指针；samplesPerChannel：每通道读取的数据；timeout mS 最长等待时间，-1持续等待直至完成，预置10000mS。只适用于一些特殊场合如与其它软件配合时使用。 | Write to AO buffer one or multi-channel raw data. The buffer is pointed by C# data type IntPtr. Parameters buf:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish,default 10000mS. Only used for very special occations such as interface with other software.
AO | AOTask | WriteSinglePoint | WriteSinglePoint(double[] buf) | Method | 向AO缓冲区写入每个通道的一个样点。参数 buf: 输出数据。缓冲区必须大于通道数。 | Write to AO buffer one sample for all channels. Parameter buf: values read. Buffer size >= number of channels. | Y
AO | AOTask | WriteSinglePoint Reload | WriteSinglePoint(double writeValue, int channel) | Method | 向AO缓冲区写入指定通道的一个样点。参数 buf: 输出数据；channel：指定通道。 | Write to AO buffer one sample for the specified channel. Parameter buf: values writen; channel: specified channel. | Y
AO | AOTask | AvaliableLenInSamples | AvaliableLenInSamples | Property | 缓冲区还可以写入的每通道点数 | AvaliableLenInSamples AO Task | Y
AO | AOTask | BufLenInSamples | BufLenInSamples | Property | 每个通道可以写入的缓冲区大小 | AO Buffer size in samples for per channels. | Y
AO | AOTask | Channels | Channels | Property | 已添加的通道列表 | A list of added AO Channels. | Y
AO | AOTask | ClockEdge | ClockEdge | Property | 时钟边沿属性。enum Rising/Falling。 | ClockEdge：enum Rising/Falling. | NA
AO | AOTask | ClockSource | ClockSource | Property | 时钟源属性。enum External/Internal。 | ClockSource: enum External/Internal. | NA
AO | AOTask | DisbaleCalibration | DisbaleCalibration | Property | 禁用校准系数 | Disalble calibration. | Y
AO | AOTask | Mode | Mode | Property | 输出模式。enum ContinuousNoWrapping/ContinuousWrapping/Finite/Single。 | Mode of AO Task. Select AO Mode from enum:ContinuousNoWrapping/ContinuousWrapping/Finite/Single. | Y
AO | AOTask | SampleClock | SampleClock | Property | 采样时钟配置。 | Sample Clock | Y
AO | AOTask | SignalExport | SignalExport | Property | 信号路由 | Signal routing | Y
AO | AOTask | Advanced | Advanced | Property | 高级属性 | Advanced properties | NA
AO | AOTask | SamplesToUpdate | SamplesToUpdate | Property | 有限点Finite模式下的输出点数 | Number of output samples in Finite mode only | Y
AO | AOTask | Trigger | Trigger | Property | 触发参数配置 | Trigger AO Task | Y
AO | AOTask | UpdateRate | UpdateRate | Property | 更新率 | UpdateRate AO Task | Y
AO | AOTask | WrappingInterval | WrappingInterval | Property | 循环输出间隔 | WrappingInterval AO Task | NA
AO | AOTask | JYXXDevice | Device | Property | 设备属性配置 | Device property set | Y
AO | AOChannel | AOChannel | AOChannel() | Constructor | 构造函数。用来建立AOChannel任务对象 | Constructor. Used to create AOChannel task object. | Y
AO | AOChannel | AOChannel | ChannelID | Property | AO通道：通道号 | AO channel ID. | Y
AO | AOChannel | AOChannel | RangeHigh | Property | AO通道：量程最大值 | High range limit an AO channel. | NA
AO | AOChannel | AOChannel | RangeLow | Property | AO通道：量程最小值 | Low range limit an AO channel. | NA
AO | AOAdvanced | AIAdvanced | AOAdvanced() | Constructor | 构造函数。用来构造AO高级设置 | Constructor to create an advanced A0 task object. | NA
AO | AOAdvanced | AIAdvanced | EnableInverseSinc | Property | 是否启用InverseSinc滤波器 | Whether enable InverseSinc filter | NA
AO | AOAdvanced | AIAdvanced | EnableInterpolation | Property | 是否使能插值功能。 | Whether enable interpolation function. | NA
AO | AOAdvanced | AIAdvanced | EnableOverflowProtection | Property | 是否启用溢出保护。 | Whether enable overflow protection. | NA
AO | AOTrigger | AOTrigger | AOTrigger() | Constructor | 构造函数。用来建立AOTrigger任务对象 | Constructor. Used to create AOTrigger task object. | Y
AO | AOTrigger | AOTrigger | Delay | Property | AO触发：触发延迟时间，单位是Tick | Number of Ticks delayed before AO triggering. | NA
AO | AOTrigger | AOTrigger | Digital | Property | AO数字触发。由两个部分组成。1 Source：硬件相关 enum；2 Edge： enum Rising/Falling。 | AO digital Trigger property has two parts. 1, Source: hardware specific enum; and 2: Edge: enum Rising/Falling. | Y
AO | AOTrigger | AOTrigger | ReTriggerCount | Property | AO触发的重触发次数。0：没有重触发。 | Number of repeat triggers. No repeat trigger when 0, or the number of repeated triggers. | NA
AO | AOTrigger | AOTrigger | Type | Property | AO触发类型：enum Immediate/Digital/Soft. | AO trigger type. enum Immediate/Digital/Soft | Y
AO | AODigitalalTrigger | AODigitalalTrigger | AODigitalalTrigger() | Constructor | 构造函数。用来建立AODigitalTrigger任务对象 | Constructor. Used to create AODigitalTrigger task object. | Y
AO | AODigitalalTrigger | AODigitalTrigger | Edge | Property | AO数字触发：触发边沿。enum Rising/Falling。 | Edge of AO digital rigger. enum Falling/Rising | Y
AO | AODigitalalTrigger | AODigitalTrigger | Source | Property | AO数字触发：外部或内部时钟源触发源。enum External/Internal。 | Clock Source of AO digital rigger. Select an external or an internal clock source from enum:External/Internal | Y
AO | AOSampleClock | AOSampleClock | AOSampleClock() | Constructor | 构造函数，创建AOSampleClock任务对象。 | Constructor to create an AOSampleClock object. | Y
AO | AOSampleClock | AOSampleClock | Source | Property | 采样时钟源，enum Internal/External | SampleClockSource，enum Internal/External | Y
AO | AOSampleClock | AOSampleClock | External | Property | 外部采样时钟 | External Sample Clock | Y
AO | AOExternalSampleClock | AOExternalSampleClock | AOExternalSampleClock() | Constructor | 构造函数，创建AOExternalSampleClock任务对象。 | Constructor to create an AOExternalSampleClock object. | Y
AO | AOExternalSampleClock | AOExternalSampleClock | Terminal | Property | 外部时钟输入端口 | External sample clock input terminal | Y
AO | AOExternalSampleClock | AOExternalSampleClock | ExpectedRate | Property | 外部时钟预期的频率 | External sample clock expected rate | Y
AO | AOSignalExport | AOSignalExport | AOSignalExport() | Constructor | 构造函数，创建AOSignalExport任务对象。 | Constructor to create an AOSignalExport object. | Y
AO | AOSignalExport | AOSignalExport | Add(AOSignalExportSource source, SignalExportDestination destination) | Method | 添加一路信号路由，将源信号路由至一个目的端口 | Add a one signal route, route the source signal to a destination terminal | Y
AO | AOSignalExport | AOSignalExport | Add(AOSignalExportSource source, List<SignalExportDestination> destinations) | Method | 添加一路信号路由，将源信号路由至多个目的端口 | Add a one signal route, route the source signal to multiple destination terminals | Y
AO | AOSignalExport | AOSignalExport | Clear( SignalExportDestination destination) | Method | 删除一路信号路由，将源信号与目的端口断开 | Removes a signal route, disconnecting the source signal from the destination terminal | Y
AO | AOSignalExport | AOSignalExport | ClearAll() | Method | 删除所有存在的信号路由 | Delete all existing signal routes | Y
DI | DITask | JYHardwareDITask | JYHardwareDITask(int slotNum) | Constructor | 构造函数。用来建立DI任务对象。slotNum板卡槽位号。硬件的唯一标志。 | Constructor. Used to create a DI task object. | Y
DI | DITask | AddChannel | AddChannel(int[] portsNum) | Method | 添加多个通道 | Add a group of channel. Parameter portsNum: an array of port numbers of added ports. | NA
DI | DITask | AddChannel Reload | AddChannel(int portNum) | Method | 添加单个通道 | Add one channel. Parameter portNum: port number to be added. | NA
DI | DITask | AddChannel Reload | AddChannel(int[] linesNum) | Method | 添加多根线 | Add a group of lines. Parameter lineNum: line number to be added. | Y
DI | DITask | AddChannel Reload | AddChannel(int lineNum) | Method | 添加单根线 | Add one line. Parameter lineNum: line number to be added. | Y
DI | DITask | AddChannel Reload | AddChannel(int portNum, int[]lineNum) | Method | 添加指定port中的多根线 | Add a group of lines in specified port . Parameter portNum: specified port. lineNum: line number to be added. | NA
DI | DITask | ReadData | ReadData(ref byte[,] dataBuffer, uint samplesToAcquire, int timeOut = 10000) | Method | 读取端口或者线的byte数据，每一列代表一个端口或者一根线。参数 dataBuffer:读回的数据缓冲; samplesToAcquire:每一端口或者线需要读取的数据长度; timeout：当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10000mS。 | Read the port or line byte data from the ports. Each column represents one port or one line. Parameters dataBuffer: data read; sampesToAcquire: number of samples per port or line to be read; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish, default 10000mS. | NA
DI | DITask | ReadData Reload | ReadData(ref byte[] dataBuffer, uint samplesToAcquire, int timeOut = 10000) | Method | 读取一个端口或者线的byte数据。参数 dataBuffer:读回的数据缓冲; samplesToAcquire:端口或者线需要读取的数据长度; timeout：当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10000mS。 | Read the byte data from one port or one line. Parameters dataBuffer: data read; sampesToAcquire: number of samples per port or one line to be read; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish, default 10000mS. | NA
DI | DITask | ReadData Reload | ReadData(ref uint[] dataBuffer,uint samplesToAcquire, int timeOut = 10000) | Method | 读取uint数据,数据中每一位代表一根线。参数 dataBuffer:读回的数据缓冲; samplesToAcquire:需要读取的数据长度; timeout：当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10000mS。 | Read uint type data, Each bit in the data represents a line. Parameters dataBuffer: uint data read; sampesToAcquire: number of samples to be read; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish, default 10000mS. | NA
DI | DITask | ReadData Reload | ReadData(ref uint[] dataBuffer,int timeOut = 10000) | Method | 读取uint数据,数据中每一位代表一根线。参数 dataBuffer:读回的数据缓冲; timeout：当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10000mS。 | Read uint type data, Each bit in the data represents a line. Parameters dataBuffer: uint data read; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish, default 10000mS. | NA
DI | DITask | ReadData Reload | ReadData(ref IntPtr dataBuffer, uint samplesToAcquire,int timeOut = 10000) | Method | 读取uint数据,数据中每一位代表一根线。缓冲区由C#数据类型IntPtr所指。参数 dataBuffer：指向缓冲区的指针；samplesPerChannel：每通道读取的数据；timeout mS 最长等待时间，-1持续等待直至完成。。只适用于一些特殊场合如与其它软件配合时使用。 | Read uint type data, Each bit in the data represents a line. The buffer is pointed by C# data type IntPtr. Parameters dataBuffer:data buffer;samplesPerChannel:number of samples per channel; timeout: timeout in mS before generating error message, timeout =-1 keep waiting until finish. Only used for very special occations such as interface with other software. | NA
DI | DITask | ReadSinglePoint | ReadSinglePoint(ref bool[] buf, int portNum) | Method | 非缓冲式读取指定端口最新的一个点，参数 buf[]读回的指定端口数据,portNum 端口号。 | Read the latest point from the specified port in unbuffered reading mode. Parameter buf[]:port value read. | NA
DI | DITask | ReadSinglePoint Reload | ReadSinglePoint(ref bool readValue, int lineNum) | Method | 非缓冲式读取指定线的最新一个点。参数 readValue读回的数据，lineNum 指定的线号。 | Read the latest point from the specified line in unbuffered reading mode . | Y
DI | DITask | ReadSinglePoint Reload | ReadSinglePoint(ref bool readValue, int portNum, int lineNum) | Method | 非缓冲式读取指定端口指定线的最新一个点。参数 readValue读回的数据，lineNum 指定的线号。 | Read the latest point from the specified line in unbuffered reading mode . | NA
DI | DITask | ReadSinglePoint Reload | ReadSinglePoint(ref bool[] buf) | Method | 非缓冲式读取所有端口或线的最新一个点。参数 buf[]读回的数据。 | Read the latest point from all ports or all lines in unbuffered reading mode. Parameter buf[]:value read. | Y
DI | DITask | RemoveChannel | RemoveChannel(int portNum) | Method | 删除指定端口号 | Delete the specified port number | NA
DI | DITask | RemoveChannel reload | RemoveChannel(int[] linesNum) | Method | 删除指定一组线号的通道 | Delete the specified a group line number | NA
DI | DITask | RemoveChannel reload | RemoveChannel(int portNum, int[] lineNum) | Method | 删除指定端口中指定线号的通道 | Delete the specified line number | NA
DI | DITask | RemoveChannel reload | RemoveChannel(int lineNum) | Method | 删除指定线号的通道 | Delete the specified line number | Y
DI | DITask | SendSoftwareTrigger | SendSoftwareTrigger() | Method | 产生一个软件触发信号。 | Generate a software trigger. | NA
DI | DITask | Start | Start() | Method | 开始DI任务 | Start DI task | Y
DI | DITask | Stop | Stop() | Method | 停止DI任务 | Stop DI task | Y
DI | DITask | AvailableSamples | AvailableSamples | Property | 缓冲区内可以读取的每个通道点数 | Number of sample points available for reading for per channe | NA
DI | DITask | BufLenInSamples | BufLenInSamples | Property | 缓冲区最大能容纳的每通道点数 | The maximum number of points per channel that the buffer can hold | NA
DI | DITask | Channels | Channels | Property | 已添加的通道列表 | List of added DI channels | Y
DI | DITask | ClockEdge | ClockEdge | Property | 时钟边沿属性。enum Rising/Falling。 | Clock edge. enum Rising/Falling. | NA
DI | DITask | ClockSource | ClockSource | Property | 时钟源属性。enum External/Internal。 | Clock Source. enum External/Internal. | NA
DI | DITask | Mode | Mode | Property | 采集模式，enum:Single/Finite/Continuous | DI Mode. enum Single/Finite/Continuous | NA
DI | DITask | SampleRate | SampleRate | Property | 采样率 | DI sampling rate per channel. | NA
DI | DITask | SamplesToAcquire | SamplesToAcquire | Property | 有限模式Finite下每一通道需要采集的点数 | Samples to be acquired per channel. Finite mode only. | NA
DI | DITask | SampleClock | SampleClock | Property | 采样时钟配置。 | Sample Clock | NA
DI | DITask | SignalExport | SignalExport | Property | 信号路由 | Signal routing | NA
DI | DITask | Trigger | Trigger | Property | 触发参数配置 | Configure DI trigger. | NA
DI | DITask | JYXXDevice | Device | Property | 设备属性配置 | Device property set | NA
DI | DITrigger | DITrigger | DITrigger() | Constructor | 构造函数，创建DITrigger任务对象。用来构造DITrigger | Constructor, used to create a DITrigger task object. | NA
DI | DITrigger | DITrigger | Mode | Property | DI触发模式属性，enum Start/Reference | DI trigger mode. enum Start/Reference. | NA
DI | DITrigger | DITrigger | Type | Property | DI触发类型属性。enum Immediate/Digital/Soft。 | DI trigger type. enum Immediate/Digital/Soft. | NA
DI | DITrigger | DITrigger | Digital | Property | DI数字触发属性。由两个部分组成。源Source：硬件相关enum; 数字触发边沿类型 Edge：enum Rising/Falling。 | DITrigger has two properties. 1 Source: hardware specific; 2 Edge: enum Rising/Falling. | NA
DI | DITrigger | DITrigger | PreTriggerSamples | Property | 触发前的预触发点数。 | Number of samples prior to triggering. | NA
DI | DITrigger | DITrigger | ReTriggerCount | Property | 重触发次数。为0时不重复触发，>0时为重复触发次数,为-1时无限重触发。 | Number of repeat triggers. No repeat trigger when 0, or the number of repeated triggers. Infinite trigger when -1. | NA
DI | DIDigitalalTrigger | DIDigitalalTrigger | DIDigitalalTrigger() | Constructor | 构造函数。用来建立DIDigitalalTrigger任务对象 | Constructor. Used to create AODigitalTrigger task object. | NA
DI | DIDigitalalTrigger | DIDigitalalTrigger | Edge | Property | 触发边沿。enum Rising/Falling。 | Edge of AO digital rigger. enum Falling/Rising | NA
DI | DIDigitalalTrigger | DIDigitalalTrigger | Source | Property | 触发源。硬件相关enum。 | Source property. hardware specific enum. | NA
DI | DIChannel | DIChannel | DIChannel() | Constructor | 构造函数，创建DIChannel任务对象。 | Constructor used to create a DIChannel task object. | Y
DI | DIChannel | DIChannel | PortNum | Property | 端口号 | port number. | NA
DI | DIChannel | DIChannel | LineNum | Property | 数字线号 | line number
DI | DISampleClock | DISampleClock | DISampleClock() | Constructor | 构造函数，创建DISampleClock任务对象。 | Constructor to create an DISampleClock object. | NA
DI | DISampleClock | DISampleClock | Source | Property | 采样时钟源，enum Internal/External | SampleClockSource，enum Internal/External | NA
DI | DISampleClock | DISampleClock | External | Property | 外部采样时钟 | External Sample Clock | NA
DI | DISampleClock | DISampleClock | ClockEdge | Property | 时钟有效边沿 | Sample clock edge | NA
DI | DIExternalSampleClock | DIExternalSampleClock | DIExternalSampleClock() | Constructor | 构造函数，创建DIExternalSampleClock任务对象。 | Constructor to create an DIExternalSampleClock object. | NA
DI | DIExternalSampleClock | DIExternalSampleClock | Terminal | Property | 外部时钟输入端口 | External sample clock input terminal | NA
DI | DIExternalSampleClock | DIExternalSampleClock | ExpectedRate | Property | 外部时钟预期的频率 | External sample clock expected rate | NA
DI | DISignalExport | DISignalExport | DISignalExport() | Constructor | 构造函数，创建DISignalExport任务对象。 | Constructor to create an DISignalExport object. | NA
DI | DISignalExport | DISignalExport | Add(DISignalExportSource source, SignalExportDestination destination) | Method | 添加一路信号路由，将源信号路由至一个目的端口 | Add a one signal route, route the source signal to a destination terminal | NA
DI | DISignalExport | DISignalExport | Add(DISignalExportSource source, List<SignalExportDestination> destinations) | Method | 添加一路信号路由，将源信号路由至多个目的端口 | Add a one signal route, route the source signal to multiple destination terminals | NA
DI | DISignalExport | DISignalExport | Clear( SignalExportDestination destination) | Method | 删除一路信号路由，将源信号与目的端口断开 | Removes a signal route, disconnecting the source signal from the destination terminal | NA
DI | DISignalExport | DISignalExport | ClearAll() | Method | 删除所有存在的信号路由 | Delete all existing signal routes | NA
DO | DOTask | JYHardwareDOTask | JYHardwareDOTask(int slotNum) | Constructor | 构造函数。用来建立DO任务对象。slotNum板卡槽位号。硬件的唯一标志。 | Constructor. Used to create a digital output DO task object. Parameter slotNum: board number, unique idendifier for the this hardware. | Y
DO | DOTask | AddChannel | AddChannel(int[] portsNum) | Method | 添加多个端口。参数：portsNum:一组端口号; | Add a group of ports. Parameter portsNum: an array of port numbers of added ports. | NA
DO | DOTask | AddChannel Reload | AddChannel(int portNum) | Method | 添加一个端口。 | Add one port. Parameter portNum: port number to be added. | NA
DO | DOTask | AddChannel Reload | AddChannel(int portNum, int[] lineNum) | Method | 添加一个端口中多根线。 | Add a group of lines in specified port . Parameter portNum: specified port. lineNum: line number to be added. | NA
DO | DOTask | AddChannel Reload | AddChannel(int[] linesNum) | Method | 添加多根线 | Add a group of lines. Parameter lineNum: line number to be added. | Y
DO | DOTask | AddChannel Reload | AddChannel(int lineNum) | Method | 添加单根线 | Add one line. Parameter lineNum: line number to be added. | Y
DO | DOTask | RemoveChannel | RemoveChannel(int portNum) | Method | 移除指定的通道。参数 portsNum：指定的通道，-1：所有通道。 | Remove specified port; remove all ports when portNum = -1. | NA
DO | DOTask | RemoveChannel reload | RemoveChannel(int[] linesNum) | Method | 删除指定一组线号的通道 | Delete the specified a group line number | NA
DO | DOTask | RemoveChannel reload | RemoveChannel(int lineNum) | Method | 删除指定线号的通道 | Delete the specified line number
DO | DOTask | RemoveChannel reload | RemoveChannel(int portNum, int[] lineNum) | Method | 删除指定端口中指定线号的通道 | Delete the specified line number | NA
DO | DOTask | SendSoftwareTrigger | SendSoftwareTrigger() | Method | 产生一个软件触发信号 | Generate a software trigger. | NA
DO | DOTask | Start | Start() | Method | 开始DO任务 | Start anDO task. | Y
DO | DOTask | Stop | Stop() | Method | 停止DO任务 | Stop an DO task. | Y
DO | DOTask | WaitUntilDone | WaitUntilDone(int timeout = -1) | Method | 等待当前任务完成。参数 timeOut: timeOut=整数N，等待N毫秒，如果没有完成，报错，timeOut=-1, 等待直至完成。 | Waiting until current task to complete. Parameter timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
DO | DOTask | WriteData | WriteData(byte[,] dataBuffer, uint samplesToUpdate, int timeOut = 10000) | Method | 写入多个端口或线的byte数据，每一列代表一个端口或线的数据。参数 dataBuffer:数据; samplesToUpdate:每个通道数据长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10000mS。 | Write multi-port or multi-line byte type data to the data buffer. Each column represents one port or one line. Parameters dataBuffer: data buffer; samplesToUpdate: number of samples for each port; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | NA
DO | DOTask | WriteData Reload | WriteData(byte[] dataBuffer, uint samplesToUpdate, int timeOut = 10000) | Method | 写入单个端口或线byte数据。参数 dataBuffer:数据; samplesToUpdate:每个通道数据长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10000mS。 | Write one port or line byte type data to the data buffer. Parameters dataBuffer: data buffer; samplesToUpdate: number of sampples for each port; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | NA
DO | DOTask | WriteData Reload | WriteData(uint[] dataBuffer, uint samplesToUpdate, int timeOut = 10000) | Method | 写入单个或多个端口unit数据,数据中每一位代表一根线。参数 dataBuffer:数据; samplesToUpdate:每个通道数据长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10002mS。 | Write one port or multi-port uint type data to the data buffer. Each bit in the data represents a line.Parameters dataBuffer: data buffer; samplesToUpdate: number of samples for each port; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | NA
DO | DOTask | WriteData Reload | WriteData(uint[] dataBuffer, int timeOut = 10000) | Method | 写入单个或多个端口unit数据,数据中每一位代表一根线。参数 dataBuffer:数据; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10002mS。 | Write one port or multi-port uint type data to the data buffer.Each bit in the data represents a line. Parameters dataBuffer: data buffer; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | NA
DO | DOTask | WriteData Reload | WriteData(IntPtr dataBuffer, uint samplesToUpdate, int timeOut = 10000) | Method | 写入单个或多个端口unit数据,数据中每一位代表一根线,缓冲区由C#数据类型IntPtr所指。参数 dataBuffer:数据; samplesToUpdate:每个通道数据长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待，预置值10002mS。 | Write one port or multi-port uint type data to the data buffer,Each bit in the data represents a line.The buffer is pointed by C# data type IntPtr.Parameters dataBuffer: data buffer; samplesToUpdate: number of samples for each port; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | NA
DO | DOTask | WriteSinglePoint | WriteSinglePoint(bool writeValue,int lineNum) | Method | 写入指定线单点值。参数 writeValue:用户缓冲区数组; lineNum：指定写入的线。 | Write a single point to a specified port. Paramters buf:data to be written; lineNum: specified line. | Y
DO | DOTask | WriteSinglePoint | WriteSinglePoint(bool writeValue,int portNum, int lineNum) | Method | 写入指定端口指定线单点值。参数 writeValue:用户缓冲区数组; lineNum：指定写入的线。 | Write a single point to a specified port. Paramters buf:data to be written; lineNum: specified line. | NA
DO | DOTask | WriteSinglePoint | WriteSinglePoint(bool[] buf,int portNum) | Method | 写入单点值。参数 buf:用户缓冲区数组; portNum：指定写入的端口。 | Write a single point to a specified port. Paramters buf:data to be written; port: specified port. | NA
DO | DOTask | WriteSinglePoint | WriteSinglePoint(bool[] buf) | Method | 写入所有端口或线的单点值，buf:用户缓冲区数组。 | Write data to all ports or lines, a single point per port or line. Paramter buf:data to be written. | Y
DO | DOTask | AvailableLenInSamples | AvailableLenInSamples | Property | 缓冲区还可以写入的最多点数。 | Number of sample points available for writing for all channels. | NA
DO | DOTask | Channels | Channels | Property | DO已添加的端口通道列表 | A list added channels of ports | Y
DO | DOTask | ClockEdge | ClockEdge | Property | 时钟边沿属性。enum Rising/Falling。 | Clock edge. enum Rising/Falling. | NA
DO | DOTask | ClockSource | ClockSource | Property | 时钟源属性。enum External/Internal。 | Clock source. enum External/Internal. | NA
DO | DOTask | Mode | Mode | Property | DO输出模式. 硬件相关enum。 | DO Mode. enum: hardware specific. | NA
DO | DOTask | SamplesToUpdate | SamplesToUpdate | Property | 每个端口有限输出点数。 | Finite output update points per port. | NA
DO | DOTask | SampleClock | SampleClock | Property | 采样时钟配置。 | Sample Clock | NA
DO | DOTask | SignalExport | SignalExport | Property | 信号路由 | Signal routing | NA
DO | DOTask | Trigger | Trigger | Property | 触发参数配置 | Configure DO trigger. | NA
DO | DOTask | UpdateRate | UpdateRate | Property | 单一端口的DO输出更新率。 | DO output update per port. | NA
DO | DOTask | JYXXDevice | Device | Property | 设备属性配置 | Device property set | NA
DO | DOChannel | DOChannel | DOChannel() | Constructor | 构造函数，创建DOChannel任务对象。 | Constructor. Used to create a DIChannel task object. | Y
DO | DOChannel | DOChannel | PortNum | Property | 端口号 | port number. | NA
DO | DOChannel | DOChannel | LineNum | Property | 数字线号 | line number | Y
DO | DOTrigger | DOTrigger | DOTrigger() | Constructor | 构造函数，创建DOTrigger任务对象。 | Constructor. Used to create a DOTrigger task object. | NA
DO | DOTrigger | DOTrigger | Digital | Property | DO数字触发属性。由两个部分组成。1 Source：硬件相关enum选; 2 Edge 数字触发边沿类型 enum:Rising/Falling。 | DOTrigger Digital property has two two parts. 1 Digital: hardware specific enum; 2 Type: enum Rising/Falling. | NA
DO | DOTrigger | DOTrigger | Type | Property | 触发类型属性，enum Immediate/Digital/Soft。 | DOTrigger Type. enum Immediate/Digital/Soft. | NA
DO | DODigitalalTrigger | DODigitalalTrigger | DODigitalalTrigger() | Constructor | 构造函数。用来建立DODigitalalTrigger任务对象 | Constructor. Used to create DODigitalTrigger task object. | NA
DO | DODigitalalTrigger | DODigitalalTrigger | Edge | Property | 数字触发边沿。enum Rising/Falling。 | Edge of DO digital trigger. enum Falling/Rising. | NA
DO | DODigitalalTrigger | DODigitalalTrigger | Source | Property | 数字触发源。硬件相关enum。 | Source of DO digital trigger. Hardware specific enum. | NA
DO | DOSampleClock | DOSampleClock | DOSampleClock() | Constructor | 构造函数，创建DOSampleClock任务对象。 | Constructor to create an DOSampleClock object. | NA
DO | DOSampleClock | DOSampleClock | Source | Property | 采样时钟源，enum Internal/External | SampleClockSource，enum Internal/External | NA
DO | DOSampleClock | DOSampleClock | External | Property | 外部采样时钟 | External Sample Clock | NA
DO | DOSampleClock | DOSampleClock | ClockEdge | Property | 时钟有效边沿 | Sample clock edge | NA
DO | DOExternalSampleClock | DOExternalSampleClock | DOExternalSampleClock() | Constructor | 构造函数，创建DOExternalSampleClock任务对象。 | Constructor to create an DOExternalSampleClock object. | NA
DO | DOExternalSampleClock | DOExternalSampleClock | Terminal | Property | 外部时钟输入端口 | External sample clock input terminal | NA
DO | DOExternalSampleClock | DOExternalSampleClock | ExpectedRate | Property | 外部时钟预期的频率 | External sample clock expected rate | NA
DO | DOSignalExport | DOSignalExport | DISignalExport() | Constructor | 构造函数，创建DOSignalExport任务对象。 | Constructor to create an DOSignalExport object. | NA
DO | DOSignalExport | DOSignalExport | Add(DOSignalExportSource source, SignalExportDestination destination) | Method | 添加一路信号路由，将源信号路由至一个目的端口 | Add a one signal route, route the source signal to a destination terminal | NA
DO | DOSignalExport | DOSignalExport | Add(DOSignalExportSource source, List<SignalExportDestination> destinations) | Method | 添加一路信号路由，将源信号路由至多个目的端口 | Add a one signal route, route the source signal to multiple destination terminals | NA
DO | DOSignalExport | DOSignalExport | Clear( SignalExportDestination destination) | Method | 删除一路信号路由，将源信号与目的端口断开 | Removes a signal route, disconnecting the source signal from the destination terminal | NA
DO | DOSignalExport | DOSignalExport | ClearAll() | Method | 删除所有存在的信号路由 | Delete all existing signal routes | NA
CI | CITask | JYHardwareCITask | JYHardwareCITask(int slotNum, int channel) | Constructor | 构造函数。用来建立CITask类,slotNum板卡槽位号。硬件的唯一标志。channel,计数器编号 | Constructor. Used to create a Counter Input CI task object. | Y
CI | CITask | ReadCounter | ReadCounter(uint[] buf, uint samplesToAcquire, int timeOut = 10000) | Method | 读取计数值。参数 buf:保存数据的缓存; samplesToAcquire:每通道读取长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read counter values. Parameters buf:buffer for counter values; samplesToAcquire: number of counter values to be read ; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadMeasure | ReadMeasure(double[] buf, uint samplesToAcquire, int timeOut = 10000) | Method | 读取测量值。参数 buf:保存数据的缓存; samplesToAcquire:每通道读取长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read measurement values. Parameters buf:buffer for the measurement values; samplesToAcquire: number of counter values to be read per channel; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadSingleCounter | ReadSingleCounter(ref uint buf) | Method | 读取单个计数值。参数 buf:保存数据的缓存 | Read a signle counter value. Parameter buf:buffer for the counter value。 | NA
CI | CITask | ReadSingleMeasure | ReadSingleMeasure(ref double buf, int timeOut = 10000) | Method | 读取单个测量值。参数 buf:保存数据的缓存; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待， 预置值10000mS。 | Read a signle measurement value. Parameters buf:buffer for the measurement value; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | NA
CI | CITask | SendSoftwareTrigger | SendSoftwareTrigger() | Method | 产生一个软件触发信号。enum Start/Reference/Pause。 | Generate a software trigger. enum Start/Reference/Pause. | NA
CI | CITask | ReadSinglePoint | ReadSinglePoint(ref uint count) | Method | 读取单个计数值。 | Read a signle counter value. Parameter buf:buffer for the counter value。 | Y
CI | CITask | ReadSinglePoint Reload | ReadSinglePoint(ref double measurement, int timeout = 10000) | Method | 读取单个测量值。 | Read a signle measurement value. Parameters measurement:buffer for the measurement value; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | Y
CI | CITask | ReadSinglePoint Reload | ReadSinglePoint(ref double measurement1, ref double measurement2, int timeout = 10000) | Method | 读取单个测量值。 | Read a signle measurement value. Parameters measurement1,measurement2:buffer for the measurement value; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate, default 10000mS. | Y
CI | CITask | ReadData | ReadData(ref uint[] bufCount, int samplesToAcquire, int timeout = 10000) | Method | 读取计数值。参数 bufCount:保存数据的缓存; samplesToAcquire:每通道读取长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read counter values. Parameters bufCount:buffer for counter values; samplesToAcquire: number of counter values to be read ; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadData Reload | ReadData(ref uint[] bufCount, int timeout = 10000) | Method | 读取计数值。参数 bufCount:保存数据的缓存; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read counter values. Parameters bufCount:buffer for counter values; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadData Reload | ReadData(ref double[] bufMeasurement, int samplesToAcquire, int timeout = 10000) | Method | 读取测量值。参数 bufMeasurement:保存数据的缓存; samplesToAcquire:每通道读取长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read measurement values. Parameters bufMeasurement:buffer for the measurement values; samplesToAcquire: number of counter values to be read per channel; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadData Reload | ReadData(ref double[] bufMeasurement, int timeout = 10000) | Method | 读取测量值。参数 bufMeasurement:保存数据的缓存; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read measurement values. Parameters bufMeasurement:buffer for the measurement values; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadData Reload | ReadData(ref double[] bufMeasurement1, ref double[] bufMeasurement2, int samplesToAcquire, int timeout = 10000) | Method | 读取测量值。参数 bufMeasurement1, bufMeasurement2 :保存数据的缓存; samplesToAcquire:每通道读取长度; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read measurement values. Parameters bufMeasurement1, bufMeasurement2:buffer for the measurement values; samplesToAcquire: number of counter values to be read per channel; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | ReadData Reload | ReadData(ref double[] bufMeasurement1, ref double[] bufMeasurement2, int timeout = 10000) | Method | 读取测量值。参数 bufMeasurement1, bufMeasurement2:保存数据的缓存; timeout当数据不足时，最多等待的时间（单位：ms），为-1时持续等待。 | Read measurement values. Parameters bufMeasurement1, bufMeasurement2:buffer for the measurement values; timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | WaitUntilDone | WaitUntilDone(int timeout = -1) | Method | 等待当前任务完成。参数 timeOut: timeOut=整数N，等待N毫秒，如果没有完成，报错，timeOut=-1, 等待直至完成。 | Waiting until current task to complete. Parameter timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CI | CITask | Start | Start() | Method | 开始CI任务 | Start CI task. | Y
CI | CITask | Stop | Stop() | Method | 停止CI任务 | Stop CI task. | Y
CI | CITask | AvailableSamples | AvailableSamples | Property | 缓冲区内还可以读取的点数 | Number of sample points available for reading . | NA
CI | CITask | TransferedSamples | TransferedSamples | Property | 已从缓冲区内读取的每通道点数 | Number of points per channel that has already been read from the buffer | NA
CI | CITask | CITimeBase | CITimeBase | Property | CI时基 | CI's time base. | NA
CI | CITask | ClockSource | ClockSource | Property | CI采样时钟源 | CI's clock source. | NA
CI | CITask | Counter | Counter | Property | 计数模式参数配置 | Configure CI Counter. | NA
CI | CITask | Measure | Measure | Property | 测量模式参数配置 | Configure CI Measure. | NA
CI | CITask | Mode | Mode | Property | CI采样模式 | CI mode | NA
CI | CITask | SampleRate | SampleRate | Property | CI采样率，仅在ClockSource为内部时钟Internal时有效。 | CI sampling rate . Valid only when ClockSource is Internal. | NA
CI | CITask | SamplesToAcquire | SamplesToAcquire | Property | CI有限采集点数 | Samples to be acquired . Finite mode only. | NA
CI | CITask | SampleClock | SampleClock | Property | 采样时钟配置。 | Sample Clock | NA
CI | CITask | SignalExport | SignalExport | Property | 信号路由 | Signal routing | NA
CI | CITask | Trigger | Trigger | Property | CI触发参数配置 | Configure CI Trigger. | NA
CI | CITask | Type | Type | Property | CI测量类型，enum:EdgeCountingt/Period/Frequency/Pulse/TwoEdgeSeparation/QuadEncoder/TwoPulseEncoder。硬件相关。 | CI Measure types. enum:EdgeCountingt/Period/Frequency/Pulse/TwoEdgeSeparation/QuadEncoder/TwoPulseEncoder. Hardware specific. | Y
CI | CITask | Sync | Sync | Property | 同步参数配置 | Sync paraments set | NA
CI | CITask | JY5211Device | Device | Property | 设备属性配置 | Device property set | Y
CI | CITrigger | CITrigger | CITrigger() | Constructor | 构造函数，创建CITrigger任务对象。 | Constructor. Used to create a CITrigger task object. | NA
CI | CITrigger | CITrigger | Digital | Property | CI数字触发属性。由两个部分组成。1 Source 硬件相关enum；2 Edge 数字触发边沿类型 enum:Rising/Falling。 | CITrigger has two parts. 1 Source: hardware specific; 2 Edge: enum Rising/Falling. | NA
CI | CITrigger | CITrigger | Type | Property | 触发类型属性，enum Immediate/Digital/Soft定。硬件相关。 | CI trigger type: enum Immediate/Digital/Soft. Hardware specific. | NA
CI | CIDigitalTrigger | CIDigitalTrigger | CIDigitalTrigger() | Constructor | 构造函数。用来建立CIDigitalalTrigger任务对象 | Constructor. Used to create AODigitalTrigger task object. | NA
CI | CIDigitalTrigger | CIDigitalTrigger | Edge | Property | 数字触发边沿。enum Rising/Falling。 | Edge of CI digital rigger. enum Falling/Rising | NA
CI | CIDigitalTrigger | CIDigitalTrigger | Source | Property | 数字触发源。硬件相关enum。 | CI trigger source. Hardware specific. | NA
CI | CISampleClock | CISampleClock | CISampleClock() | Constructor | 构造函数，创建CISampleClock任务对象。 | Constructor to create an CISampleClock object. | NA
CI | CISampleClock | CISampleClock | Source | Property | 采样时钟源，enum Internal/External/Implicit | SampleClockSource，enum Internal/External/Implicit | NA
CI | CISampleClock | CISampleClock | External | Property | 外部采样时钟 | External Sample Clock | NA
CI | CISampleClock | CISampleClock | Internal | Property | 内部采样时钟 | Internal Sample Clock | NA
CI | CISampleClock | CISampleClock | Implicit | Property | 隐式时钟 | Implicit Sample Clock | NA
CI | CIExternalSampleClock | CIExternalSampleClock | CIExternalSampleClock() | Constructor | 构造函数，创建CIExternalSampleClock任务对象。 | Constructor to create an CIExternalSampleClock object. | NA
CI | CIExternalSampleClock | CIExternalSampleClock | Terminal | Property | 外部时钟输入端口 | External sample clock input terminal | NA
CI | CIExternalSampleClock | CIExternalSampleClock | ExpectedRate | Property | 外部时钟预期的频率 | External sample clock expected rate | NA
CI | CIInternalSampleClock | CIInternalSampleClock | CIInternalSampleClock() | Constructor | 构造函数，创建CIInternalSampleClock任务对象。 | Constructor to create an CIInternalSampleClock object. | NA
CI | CIInternalSampleClock | CIInternalSampleClock | Rate | Property | 内部时钟采样率 | The rate of the internal sample clock | NA
CI | CIImplicitSampleClock | CIImplicitSampleClock | CIImplicitSampleClock() | Constructor | 构造函数，创建CIImplicitSampleClock任务对象。 | Constructor to create an CIImplicitSampleClock object. | NA
CI | CIImplicitSampleClock | CIImplicitSampleClock | ExpectedRate | Property | 隐式时钟预期的频率 | Implicit sample clock expected rate | NA
CI | CISignalExport | CISignalExport | CISignalExport() | Constructor | 构造函数，创建CISignalExport任务对象。 | Constructor to create an CISignalExport object. | NA
CI | CISignalExport | CISignalExport | Add(CISignalExportSource source, SignalExportDestination destination) | Method | 添加一路信号路由，将源信号路由至一个目的端口 | Add a one signal route, route the source signal to a destination terminal | NA
CI | CISignalExport | CISignalExport | Add(CISignalExportSource source, List<SignalExportDestination> destinations) | Method | 添加一路信号路由，将源信号路由至多个目的端口 | Add a one signal route, route the source signal to multiple destination terminals | NA
CI | CISignalExport | CISignalExport | Clear( SignalExportDestination destination) | Method | 删除一路信号路由，将源信号与目的端口断开 | Removes a signal route, disconnecting the source signal from the destination terminal | NA
CI | CISignalExport | CISignalExport | ClearAll() | Method | 删除所有存在的信号路由 | Delete all existing signal routes | NA
CI | Counter | Counter | Counter() | Constructor | 构造函数，创建Counter任务对象。 | Constructor. Used to create AO Counter task object. | NA
CI | Counter | Counter | ClockEdge | Property | 计数模式下的时钟有限边沿，Rising/Falling | Active edge of clock in counter mode | NA
CI | Counter | Counter | ClockSource | Property | 计数模式下的时钟源。enum External/Internal。 | clock source in counter mode | NA
CI | Counter | Counter | Direction | Property | 计数方向。enum Up/Down. Up: 计数值增加; Down: 计数值减小。External:外部信号控制计数方向，高电平时增计数，低电平时减计数 | Counter direction. enum Up/Down. Up: Counter increases; Down: counter decreases. External: External signal control counter direction. If External signal is highlevel counter increases, or lowlevel counter decreases. | NA
CI | Counter | Counter | DirectionPolarity | Property | 计数方向的有效电平 | Active Level of Counter direction | NA
CI | Counter | Counter | GatePolarity | Property | 门信号的有效电平 | Active Level of Gate Signal | NA
CI | Counter | Counter | GateSource | Property | 门信号源 | Gate Signal Source | NA
CI | Counter | Counter | InitialCount | Property | 计数器初始值。 | Counter initial value. | NA
CI | Measure | Measure | Measure() | Constructor | 构造函数，创建Measure任务对象。 | Constructor. Used to create AO Measure task object. | NA
CI | Measure | Measure | ClockEdge | Property | 测量模式下的时钟有限边沿，Rising/Falling | Active edge of clock in measure mode. enum Rising/Falling. | NA
CI | Measure | Measure | ClockSource | Property | 测量模式下的时钟源。enum External/Internal。 | clock source in measure mode.enum External/Internal。 | NA
CI | Measure | Measure | EdgeType | Property | 测量双边沿间隔的类型 | Edge types used in the measure. enum FirstToSecond/SecondToFirst. FirstToSecond:Gate rising edge to AUX rising edge; SecondToFirst:AUX rising edge to Gate rising edge. | NA
CI | Measure | Measure | ExternalClockRate | Property | 外部时钟源的频率 | external clock rare | NA
CI | Measure | Measure | PulseType | Property | 测量脉冲的类型 | Pulse types used in the measure. | NA
CI | Measure | Measure | Type | Property | 测量类型。SinglePeriodMSR/SinglePulseWidthMSR/EdgeSeparationMSR | measure type。 enum：SinglePeriodMSR/SinglePulseWidthMSR/EdgeSeparationMSR | NA
CI | EdgeCounting | EdgeCounting | EdgeCounting() | Constructor | 构造函数，创建EdgeCounting任务对象。 | Constructor to create an EdgeCounting object. | Y
CI | EdgeCounting | EdgeCounting | InitialCount | Property | 计数器初始值。 | Counter initial value. | Y
CI | EdgeCounting | EdgeCounting | ActiveEdge | Property | 信号有效边沿 | Signal active edge | NA
CI | EdgeCounting | EdgeCounting | Direction | Property | 计数方向。enum Up/Down. Up: 计数值增加; Down: 计数值减小。External:外部信号控制计数方向，高电平时增计数，低电平时减计数 | Counter direction. enum Up/Down. Up: Counter increases; Down: counter decreases. External: External signal control counter direction. If External signal is highlevel counter increases, or lowlevel counter decreases. | Y
CI | EdgeCounting | EdgeCounting | OutEvent | Property | 边沿计数输出事件。当计数值等于指定阈值时输出一个脉冲 | Edge counting output event. A pulse will be generated when the count value reaches specified threshold | Y
CI | EdgeCounting | EdgeCounting | Pause | Property | 暂停触发配置 | Pause Trigger Configuration | Y
CI | EdgeCounting | EdgeCounting | InputTerminal | Property | 被测信号的输入端口 | Input Terminal of Signal to Measure | NA
CI | EdgeCounting | EdgeCounting | DirTerminal | Property | 外部方向控制信号的输入端口 | Terminal of External Count Diretion Control | NA
CI | EdgeCounting | EdgeCounting | PauseTrigger() | Constructor | 构造函数，创建PauseTrigger任务对象。 | Constructor to create an PauseTrigger object. | Y
CI | EdgeCounting | EdgeCounting | Terminal | Property | 暂停触发的输入端口 | Pause Trigger Input Terminal | NA
CI | EdgeCounting | EdgeCounting | ActivePolarity | Property | 暂停触发有效极性 | Pause trigger active polarity | Y
CI | FrequencyMeas | FrequencyMeas | FrequencyMeas() | Constructor | 构造函数，创建FrequencyMeas任务对象。 | Constructor to create an FrequencyMeas object. | Y
CI | FrequencyMeas | FrequencyMeas | Timebase | Property | 测量时基 | Measure time base. | NA
CI | FrequencyMeas | FrequencyMeas | InputTerminal | Property | 被测信号的输入端口 | Input Terminal of Signal to Measure | NA
CI | PeriodMeas | PeriodMeas | PeriodMeas() | Constructor | 构造函数，创建PeriodMeas任务对象。 | Constructor to create an PeriodMeas object. | Y
CI | PeriodMeas | PeriodMeas | Timebase | Property | 测量时基 | Measure time base. | NA
CI | PeriodMeas | PeriodMeas | InputTerminal | Property | 被测信号的输入端口 | Input Terminal of Signal to Measure | NA
CI | PulseMeas | PulseMeas | PulseMeas() | Constructor | 构造函数，创建PulseMeas任务对象。 | Constructor to create an PulseMeas object. | Y
CI | PulseMeas | PulseMeas | Timebase | Property | 测量时基 | Measure time base. | NA
CI | PulseMeas | PulseMeas | InputTerminal | Property | 被测信号的输入端口 | Input Terminal of Signal to Measure | NA
CI | TwoEdgeSeparation | TwoEdgeSeparation | TwoEdgeSeparation() | Constructor | 构造函数，创建TwoEdgeSeparation任务对象。 | Constructor to create an TwoEdgeSeparation object. | Y
CI | TwoEdgeSeparation | TwoEdgeSeparation | Timebase | Property | 测量时基 | Measure time base. | NA
CI | TwoEdgeSeparation | TwoEdgeSeparation | FirstInputTerminal | Property | 被测信号的输入端口 | Input Terminal of Signal to Measure | NA
CI | TwoEdgeSeparation | TwoEdgeSeparation | SecondInputTerminal | Property | 被测信号的输入端口 | Input Terminal of Signal to Measure | NA
CI | QuadEncoder | QuadEncoder | QuadEncoder() | Constructor | 构造函数，创建QuadEncoder任务对象。 | Constructor to create an QuadEncoder object. | Y
CI | QuadEncoder | QuadEncoder | EncodingType | Property | 编码类型，enum X1/X2/X4 | Encoding type enum X1/X2/X4 | Y
CI | QuadEncoder | QuadEncoder | ZReloadEnabled | Property | Z信号复位使能 | Z Signal reload enabled | Y
CI | QuadEncoder | QuadEncoder | InitialCount | Property | 计数器初始值。 | Counter initial value. | Y
CI | QuadEncoder | QuadEncoder | AInputTerminal | Property | 信号A的输入端口 | Input Terminal of Signal A | NA
CI | QuadEncoder | QuadEncoder | BInputTerminal | Property | 信号B的输入端口 | Input Terminal of Signal B | NA
CI | QuadEncoder | QuadEncoder | ZInputTerminal | Property | 信号Z的输入端口 | Input Terminal of Signal Z | NA
CI | TwoPulseEncoder | TwoPulseEncoder | TwoPulseEncoder() | Constructor | 构造函数，创建TwoPulseEncoder任务对象。 | Constructor to create anTwoPulseEncoder object. | NA
CI | TwoPulseEncoder | TwoPulseEncoder | InitialCount | Property | 计数器初始值。 | Counter initial value. | NA
CI | TwoPulseEncoder | TwoPulseEncoder | AInputTerminal | Property | 信号A的输入端口 | Input Terminal of Signal A | NA
CI | TwoPulseEncoder | TwoPulseEncoder | BInputTerminal | Property | 信号B的输入端口 | Input Terminal of Signal B | NA
CI | CITimebase | CITimebase | CITimebase() | Constructor | 构造函数，创建CITimebase任务对象。 | Constructor to create an CITimebase object. | NA
CI | CITimebase | CITimebase | Source | Property | 时基源 | timebase | NA
CI | CITimebase | CITimebase | External | Property | 外部时基配置 | External Timebase Configuration | NA
CI | CIExternalTimebase | CIExternalTimebase | CIExternalTimebase() | Constructor | 构造函数，创建CIExternalTimebase任务对象。 | Constructor to create an CIExternalTimebase object. | NA
CI | CIExternalTimebase | CIExternalTimebase | Terminal | Property | 外部时基的输入端口 | External Timebase Input Terminal | NA
CI | CIExternalTimebase | CIExternalTimebase | Frequency | Property | 外部时基的频率 | Frequency of The External Timebase | NA
CO | COTask | JYHardwareCOTask | JYHardwareCOTask(int slotNum, int channel) | Constructor | 构造函数。用来建立COTask类,slotNum板卡槽位号。硬件的唯一标志。channel,计数器编号 | Constructor. Used to create a counter output CO task object. | Y
CO | COTask | ApplyParam | ApplyParam() | Method | 实时更新DutyCycle及Frequency参数。 | Update DutyCycle and Frequency parameters in real time. | NA
CO | COTask | SendSoftwareTrigger | SendSoftwareTrigger() | Method | 产生一个软件触发信号 | Generate a software trigger. | NA
CO | COTask | Start | Start() | Method | 开始CO任务 | Start a CO task. | Y
CO | COTask | Stop | Stop() | Method | 停止CO任务 | Stop a CO task. | Y
CO | COTask | WriteSinglePoint | WriteSinglePoint(COPulse pulse) | Method | 写入单个点 | Write one point | Y
CO | COTask | WriteData | WriteData(COPulse[] pulsesGroup, int samplesToWrite, int timeout = 10000) | Method | 写入一组数据 | Write pulse parameters to be output to the buffer | NA
CO | COTask | WaitUntilDone | WaitUntilDone(int timeout = -1) | Method | 等待当前任务完成。参数 timeOut: timeOut=整数N，等待N毫秒，如果没有完成，报错，timeOut=-1, 等待直至完成。 | Waiting until current task to complete. Parameter timeout = -1, wait until done, other positive number N, wait N mS. If not done, a timeout error generate. | NA
CO | COTask | Clock | Clock | Property | CO时钟信号配置,可配置时钟源Interna/External和时钟有限边沿Rising/Falling | Configure CO Clock. | NA
CO | COTask | COTimeBase | TimeBase | Property | CO时基 | CO timebase. | NA
CO | COTask | Gate | Gate | Property | CO门信号配置，可配置门信号源Interna/External和门信号有限电平HighActive/LowActive | Configure CO Gate signal. | NA
CO | COTask | IdleState | IdleState | Property | 输出开始前和停止后的空闲状态 | Idle state when pulse output before start and after stop | Y
CO | COTask | InitialDelay | InitialDelay | Property | 输出脉冲的初始延迟, 以Tick为单位 | Initial CO output pulse delay in number of Ticks. | NA
CO | COTask | OutputTerminal | OutputTerminal | Property | 脉冲输出端口 | Pulse Output Terminal | NA
CO | COTask | Sync | Sync | Property | 同步参数配置 | Sync paraments set | NA
CO | COTask | Device | Device | Property | 设备属性配置 | Device property set | NA
CO | COTask | SamplesToUpdate | SamplesToUpdate | Property | 要写入到Buffer的Sample数 | The Sample number to be written to the buffer | NA
CO | COTask | AvailableBufferLength | AvailableBufferLength | Property | 板载缓冲区当前可写入的样点数（每通道） | The number of samples that can be written into the onboard buffer (per channel). | NA
CO | COTask | TransferedSamples | TransferedSamples | Property | 已从本地缓冲区传出的样点数（每通道） | The number of samples that have been transfered from the local buffer (per channel) | NA
CO | COTask | TransferedPulses | TransferedPulses | Property | 当前配置下已发送的脉冲的数量 | The number of pulses transfered under the current configuration | NA
CO | COTask | Mode | Mode | Property | CO输出模式 | CO Mode. | NA
CO | COTask | Pulse | Pulse | Property | 输出脉冲配置 | Configure CO Output pulse. | Y
CO | COTask | Trigger | Trigger | Property | CO触发参数配置 | Configure CO trigger. | NA
CO | COTimebase | COTimebase | COTimebase() | Constructor | 构造函数，创建COTimebase任务对象。 | Constructor to create an COTimebase object. | NA
CO | COTimebase | COTimebase | Source | Property | 时基源 | timebase | NA
CO | COTimebase | COTimebase | External | Property | 外部时基配置 | External Timebase Configuration | NA
CO | COExternalTimebase | COExternalTimebase | COExternalTimebase() | Constructor | 构造函数，创建COExternalTimebase任务对象。 | Constructor to create an COExternalTimebase object. | NA
CO | COExternalTimebase | COExternalTimebase | Terminal | Property | 外部时基的输入端口 | External Timebase Input Terminal | NA
CO | COExternalTimebase | COExternalTimebase | Frequency | Property | 外部时基的频率 | Frequency of The External Timebase | NA
CO | COPulse | COPulse | COPulse() | Constructor | 构造函数，创建COPulse任务对象。 | Constructor. Used to create a CO pulse task object. | NA
CO | COPulse | COPulse | Count | Property | 有限脉冲输出个数 | Pulse count in finite mode | NA
CO | COPulse | COPulse | DutyCycleFrequency | Property | 按占空比和频率设置脉冲属性。由两个部分组成。1，DutyCycle 输出脉冲的占空比; 2，Frequency 输出脉冲的频率。 | Configure CO output pulse according to its DutyCycle and Freqency. It has two parts. 1 DutyCycle; 2 Frequency. | NA
CO | COPulse | COPulse | InitialDelay | Property | 输出脉冲的初始延迟, 以Tick为单位 | Initial CO output pulse delay in number of Ticks. | NA
CO | COPulse | COPulse | Tick | Property | 使用Tick设置脉冲属性。由两个属性组成。1.High 脉冲高电平Tick; 2，Low 脉冲低电平Tick。 | Configure CO output pulse according to its high/low tick voltage level. It has two parts. 1: High Tick level; 2: Low Tick Level. | NA
CO | COPulse | COPulse | Time | Property | 使用高低电平时间设置脉冲属性。由两个属性组成。1.High 脉冲高电平时间; 2，Low 脉冲低电平时间。 | Configure CO output pulse according to duration of its high/low time voltage levels. It has two parts. 1 High voltage level; 2 Low voltage level. | NA
CO | COPulse | COPulse | Type | Property | 脉冲设置类型。enum DutyCycleFrequencyenum/HighLowTime/HighLowTick。 | set pulse type. enum DutyCycleFrequencyenum/HighLowTime/HighLowTick. | NA
CO | COPulse | COPulse | COPulse(COPulseType type, double param1, double param2, int count) | Method | 脉冲参数设置。 | Pulse paraments set | NA
CO | COTrigger | COTrigger | COTrigger() | Constructor | 构造函数，创建COTrigger任务对象。 | Constructor. Used to create a COTrigger task object. | NA
CO | COTrigger | COTrigger | Digital | Property | CO数字触发由两部分组成。1 Source：硬件相关 enum；2 Edge： enum Rising/Falling。 | COTrigger Digital property has two parts. 1 Source: hardware specific enum; and 2: Edge: enum Rising/Falling. | NA
CO | COTrigger | COTrigger | Type | Property | CO触发类型。enum :Immediate/Digital/Soft | COTrigger Type: enum Immediate/Digital/Soft. | NA
CO | CODigitalTrigger | CODigitalTrigger | CODigitalTriggerr() | Constructor | 构造函数。用来建立CODigitalalTrigger任务对象 | Constructor. Used to create CODigitalTrigger task object. | NA
CO | CODigitalTrigger | CODigitalTrigger | Edge | Property | 触发边沿。enum Rising/Falling。 | Edge of CO digital rigger. enum Falling/Rising. | NA
CO | CODigitalTrigger | CODigitalTrigger | Source | Property | 触发源。enum:PFI_0~PFI_15,PXITrigger0~7 | Source of CO digital rigger. Hardware specific enum. | NA
CO | DutyCycleFrequency | DutyCycleFrequency | DutyCycleFrequency() | Constructor | 构造函数，创建DutyCycleFrequency任务对象。 | Constructor. Used to create a DutyCycleFrequency task object. | Y
CO | DutyCycleFrequency | DutyCycleFrequency | DutyCycle | Property | 输出脉冲的占空比 | Dutycyle of CO output pulse. | Y
CO | DutyCycleFrequency | DutyCycleFrequency | Frequency | Property | 输出脉冲的频率 | Frequency of CO output pulse. | Y
CO | Time | Time | Time() | Constructor | 构造函数，创建Time任务对象。 | Constructor. Used to create CO's timebase task object. | Y
CO | Time | Time | High | Property | 输出脉冲的高电平时间 | High voltage duration of CO output pulse. | Y
CO | Time | Time | Low | Property | 输出脉冲的低电平时间 | Low voltage duration of CO output pulse. | Y
CO | Tick | Tick | Tick() | Constructor | 构造函数，创建Tick任务对象。 | Constructor. Used to create a CO Tick task object. | Y
CO | Tick | Tick | High | Property | 输出脉冲的高电平Tick数 | Number of high volage ticks of CO output puse. | Y
CO | Tick | Tick | Low | Property | 输出脉冲的低电平Tick数 | Number of low volage ticks of CO output puse. | Y
CO | COClock | COClock | COClock() | Constructor | 构造函数用来构造CO Clock任务对象 | Constructor. Used to create COClock task object. | NA
CO | COClock | COClock | Edge | Property | 时钟边沿属性。enum Rising/Falling。 | Active edge of Gate Signal. enum Rising/Falling | NA
CO | COClock | COClock | Souce | Property | 时钟源属性。enum External/Internal。 | Clock source. enum External/Internal. | NA
CO | COGate | COGate | COGate | Constructor | 构造函数用来构造CO Gate任务对象 | Constructor. Used to create COGate task object. | NA
CO | COGate | COGate | Polarity | Property | 门信号有效电平HighActive/LowActive | Active Level of Gate Signal. enum HighActive/LowActive | NA
CO | COGate | COGate | Source | Property | 门信号源External/Internal | Gate Signal Source.enum External/Internal. | NA
Device | Device | Device | Device() | Constructor | 构造函数，创建Device任务对象。 | Constructor to create an Device object. | NA
Device | Device | Device | ReferenceClock | Property | 参考时钟配置 | Reference Clock Configuration | NA
Device | Device | Device | PFI | Property | PFI 配置 | PFI config | NA
Device | ReferenceClock | ReferenceClock | ReferenceClock() | Constructor | 构造函数，创建ReferenceClock任务对象。 | Constructor to create an ReferenceClock object. | NA
Device | ReferenceClock | ReferenceClock | Source | Property | 参考时钟源 | Reference Clock Source | NA
Device | ReferenceClock | ReferenceClock | External | Property | 外部参考时钟源配置 | External Reference Clock Source | NA
Device | ExternalReferenceClock | ExternalReferenceClock | ExternalReferenceClock() | Constructor | 构造函数，创建ExternalReferenceClock任务对象。 | Constructor to create an ExternalReferenceClock object. | NA
Device | ExternalReferenceClock | ExternalReferenceClock | Terminal | Property | 外部参考时钟输入端口 | External Reference Clcok Input Terminal | NA
Device | ExternalReferenceClock | ExternalReferenceClock | Frequency | Property | 外部参考时钟频率 | Frequency of External Reference Clcok | NA
Device | PFI | PFI | PFI() | Constructor | 构造函数，创建PFI任务对象。 | Constructor to create an PFI object. | NA
Device | PFI | PowerLevel | LogicLevel | Property | PFI端口的输入/输出高的电平标准 | Input/Output high voltage of PFIs | NA
Device | PFI | PFIFilter | Filter | Property | PFI滤波器配置 | PFI Filter Setting | NA
Device | PFIFilter | PFIFilter | PFIFilter() | Constructor | 构造函数，创建PFIFilter任务对象。 | Constructor to create an PFIFilter object. | NA
Device | PFIFilter | Enable | Enable(IOTerminal terminal, int minPulseWidth) | Method | 启用指定端口上的滤波器并设置最小脉宽 | Enable the filter on the specified terminal and set the minimum pulse width | NA
Device | PFIFilter | Enable Reload | Enable(IOTerminal[] terminal, int minPulseWidth) | Method | 启用一组端口上的滤波器并设置最小脉宽 | Enable the filter on a group of terminals and set the minimum pulse width | NA
Device | PFIFilter | Enable Reload | Enable(IOTerminal[] terminal, int[] minPulseWidth) | Method | 启用一组端口上的滤波器并分别设置最小脉宽 | Enable the filter on a group of terminals and set the minimum pulse width separately | NA
Device | PFIFilter | Disable | Disable(IOTerminal terminal) | Method | 禁用指定端口上的滤波器 | Disable the filter on the specified terminal | NA
Device | PFIFilter | Disable Reload | Disable(IOTerminal[] terminal) | Method | 禁用一组端口上的滤波器 | Disable the filter on a group terminal | NA
Device | PFIFilter | DisableAll | DisableAll() | Method | 禁用所有端口上的滤波器 | Disable the filter on all the terminals | NA